- List of Reagents examined in this report:

| Reagents | Lot Number |
| --- | --- |
| 10 x TE | 37226 |

- The associated equipment:

Tecan

- The acceptance Criteria for the results being run on that particular equipment

- Target volume:  is 105 ng/ul and the acceptance range is +/- 10 ng/ul of the target (ie) 95 ng/ul to 115 ng/ul.

- A260 concentration: Target is 60 ng/ul and the acceptance range is +10 % of the Target A260 concentration (ie) 54 ng/ul to 66 ng/ul.

- A260/A280 absorbance ratio: 1.60 to 2.04

- Results of the quarterly validation

|  | Mean Volume (ng/ul) |  | A260 concentration (ng/ul) |  | A260/A280 absorbance ratios |  |
| --- | --- | --- | --- | --- | --- | --- |
| Tecan | Mean | Standard deviation | Mean | Standard deviation | Mean | Standard deviation |
|  | 103.19 | 2.72 | 57.93 | 4.52 | 1.77 | 0.06 |

- Do the quarterly validation results fall within acceptance criteria?

- Yes

  - The mean value of A260 concentration (ng/ul) is within the acceptance level (> 54 ng/ul)

  - The mean value of A260/A280 absorbance ratio within the acceptance range

  - The mean volume is also in the acceptance limit.

- Can the reagents mentioned above be signed off as validated?

  - Yes