List of Reagents examined in this report:

| Reagent | Lot Number |
| --- | --- |
| Cell Lysis | 0000481464 |
| Binding Buffer | 0000487523 |
| Prepared Wash Buffer | 0000485667 |
| 50% Ethanol | 0000121067 |
| Reliaprep Reagent | 0000476709 |
| Alkaline Protease | 0000334807 |
| Pro K | 0000488401 |
| Nuclease Free Water (invitrogen) | 2107460 |
| Nuclease Free Water (invitrogen) | 2011411 |
| 10 x TE | 164977 |
| Sodium Hydroxide | 20J304019 |
| 96% Ethanol | 1923760 |

The associated equipment:

- HSM 1

The acceptance Criteria for the results being run on that particular equipment

- Mean A260 concentration (ng/ul) > 15 ng/ul

- A260/A280 ratio between 1.60 – 2.04

Results of the quarterly validation

|  | A260 concentration (ng/ul) |  | A260/A280 absorbance ratios |  |
| --- | --- | --- | --- | --- |
| Tecan-HSM | Mean | Standard deviation | Mean | Standard deviation |
| 1 | 136.51 | 93.92 | 1.66 | 0.17 |

Do the quarterly validation results fall within acceptance criteria?

- Yes

  - The mean value of A260 concentration (ng/ul) is > 15 ng/ul and the validation is passed

  - The mean value of A260/A280 absorbance ratio is also within the acceptance range

The Negative control values are +/- 3 standard deviation for that batch tested.

Can the reagents mentioned above be signed off as validated?

Yes