- List of Reagents examined in this report:

| Reagents | Lot Number |
| --- | --- |
| RNAse Free Water | 169038013 |
| RNAse Free Water | 169016239 |
| BM4 | 169038304 |
| BM4 | 169021219 |
| BM3 | 169037825 |
| BM1 | 169016329 |
| BM2 | 172011473 |
| Pro K | 169045035 |
| BR5 | 169037198 |
| Paxgene RNA Spin Column | 169047761 |
| Paxgene RNA Shredder | 169049789 |
| Rotor Adaptors | CH-2143574 |
| Ethanol 96% | 11911938 |
| Ethanol 96% | 1923760 |

- The associated equipment:

QIAcube 1

- The acceptance Criteria for the results being run on that particular equipment

  - Mean A260 concentration (ng/ul) > 50 ng/ul

  - A260/A280 ratio between 1.9 – 2.1

- Results of the quarterly validation

|  | A260 concentration (ng/ul) |  | A260/A280 absorbance ratios |  |
| --- | --- | --- | --- | --- |
| Qiacube | Mean | Standard deviation | Mean | Standard deviation |
| 1 | 56.62 | 6.86 | 2.09 | 0.01 |

- Do the quarterly validation results fall within acceptance criteria?

- Yes

  - The mean value of A260 concentration (ng/ul) is > 50 ng/ul and the validation is passed

  - The mean value of A260/A280 absorbance ratio within the acceptance range

- Can the reagents mentioned above be signed off as validated?

  - Yes

Negative control value is nearer to the mean for that batch tested.

Can the reagents mentioned above be signed off as validated?

Yes